## Slide 1: Uncertainty Essentials
Measurement Uncertainty in Testing: What it is and what its for

## Slide 2: Uncertainty Essentials
Definition of Measurement Uncertainty
  Expression of Measurement Uncertainty
17025:2017 & 15189:2012 Requirements
  Use in selection of methods
  Use in reporting
Statistics, Units and Calculations
  Standard deviation
  Transformation of non-normal data
  Harmonisation of scales & units
Approaches to Estimation of Measurement Uncertainty
  Metrological approach
  Single Within-Laboratory Approach
Qualitative Results
Guidance Documents

## Slide 3: Definition of Measurement Uncertainty
A parameter associated with the result of a measurement, that characterises the dispersion of the values that could reasonably be attributed to the measurand
  ISO/IEC Guide 98 GUM
Uncertainty is a property of a test result
Preferred form of reporting:
  Result: (x ± U) (units)

## Slide 4: Reference to Measurement Uncertainty
Knowledge of method performance often contributes to estimates of uncertainty
Reference is often made to
‘uncertainty of a method’
Uncertainty must always be associated with a test result
  or range of test results
X

## Slide 5: 17025:2017
7.6.1 Laboratories shall identify the contributions to measurement uncertainty. When evaluating measurement uncertainty, all contributions that are of significance, including those arising from sampling, shall be taken into account using appropriate methods of analysis.
7.6.3 A laboratory performing testing shall evaluate measurement uncertainty. Where the test method precludes rigorous evaluation of measurement uncertainty, an estimation shall be made based on an understanding of the theoretical principles or practical experience of the performance of the method.

## Slide 6: 17025:2017
6.4.5 Equipment performance specifications
7.2 Selection, Verification and validation of methods
7.8.3 Test Reports: where applicable, the measurement uncertainty presented in the same unit as that of the measurand or in a term relative to the measurand (e.g. percent) when:
  it is relevant to the validity or application of the test results;
  a customer's instruction so requires, or
  the measurement uncertainty affects conformity to a specification limit;

## Slide 7: 17025:2017
NOTES
In those cases where a well-recognized test method specifies limits to the values of the major sources of measurement uncertainty and specifies the form of presentation of the calculated results, the laboratory is considered to have satisfied 7.6.3 by following the test method and reporting instructions.
For a particular method where the measurement uncertainty of the results has been established and verified, there is no need to evaluate measurement uncertainty for each result if the laboratory can demonstrate that the identified critical influencing factors are under control.

## Slide 8: ISO 15189:2012
“The laboratory shall determine measurement uncertainty for each measurement procedure in the examination phases used to report measured quantity values on patients’ samples”
“The laboratory shall define the performance requirements for the measurement uncertainty of each measurement procedure and regularly review estimates of measurement uncertainty .”
“upon request, the laboratory should make its estimates of measurement uncertainty available to laboratory users.”

## Slide 9: Standard Uncertainty
Uncertainties are usually quantified as standard deviations
Standard uncertainty, u
Expanded uncertainty, U = k.u
  k = coverage factor, usually 2
  The reported uncertainty is an expanded uncertainty calculated using a coverage factor of 2 which gives a level of confidence of approximately 95 %
Estimation of uncertainty for test methods utilizes the arithmetic of standard deviations.

## Slide 10: Transformation & Harmonisation
Use of standard deviations implies that data conforms to normal distribution
Data from some types of measurement may need to be transformed before they can analysed
  E.g. microbiological counts – Log10
  Results from quantitative PCR measurements - Log10
Data from calibration certificates and other sources may need to be converted to standard deviations

## Slide 11: Transforming Ranges to Standard Deviations
Source | Transformation
Precision | sr and sR used directly
Bias | sB estimated from validation data – see later
Expanded uncertainties from calibration certificates | Divide by 2 (or k)
Tolerance limits stated as ±x | Rectangular distribution: divide by √3
Tolerance limits stated as ±x, but where manufacturer aims to be close to specification | Triangular distribution: divide by √6
Constant weight limits | Rectangular distribution: divide by √3

## Slide 12: Approaches to Estimating Uncertainty
Modelling Approach (“bottom-up”)
  Used extensively for metrological testing methods but rarely successfully applied to complete chemical or microbiological methods
  Refs: Eurachem 2012
Single Laboratory Validation Approach (“top down”)
  Applied extensively for chemical or microbiological methods, relying principally on estimates of reproducibility and bias
  Internal estimates of reproducibility and bias, and/or use of PT results to estimate bias
  Ref: Eurachem 2012, Nordtest 2004, Eurolab 2007

## Slide 13: Single Within-Laboratory Validation
Measurement accuracy = precision + trueness
Measurement uncertainty = within-laboratory reproducibility + uncertainty on bias
Where u = standard uncertainty, sR is reproducibility of method, b is bias of method at level of test result.

## Slide 14: Uncertainty in Qualitative Results
Uncertainty in qualitative results not addressed in general guidance documents
  Yes/no, presence/absence, pass/fail
Some published papers on estimation using complex statistical techniques e.g. Bayes theorem.
  Require high expertise & software
Some aspects of verification may implicitly include some sources of uncertainty
  Setting of cut-off values

## Slide 15: Uncertainty in Qualitative Results
Where results are expressed in qualitative terms, there is no expectation of a quantitative estimate of measurement uncertainty
  Unless a prior quantitative test result produced
Factors affecting the test result need to be identified and shown to be appropriately controlled.

## Slide 16
The phase of the moon?
Analytical Methods Committee, AMCTB No 61

## Slide 17: Guidance Documents
ISO/IEC Guide 98-1:2009 Uncertainty of measurement
  Part 1: Introduction to the expression of uncertainty in measurement
  Part 3: Guide to the expression of uncertainty in measurement (GUM:1995)
EURACHEM / CITAC Guide CG 4 Quantifying Uncertainty in Analytical Measurement, Ed 3, 2012
Eurolab Technical Report No 1/2007. Measurement Uncertainty Revisited: Alternative Approaches to Uncertainty.
Handbook for calculation of measurement uncertainty in environmental laboratories (NT TR 537 - Edition 3.1)
EURACHEM / CITAC Guide Measurement Uncertainty Arising From Sampling - A Guide To Methods And Approaches, Ed 1, 2007
ISO 5725 Accuracy (Trueness and Precision) of Measurement Methods and Results (6 parts)
The “Phase-of-the-Moon” Paradox in Uncertainty Estimation. Analytical Methods Committee, AMCTB No 61

## Slide 18: Uncertainty Essentials
Definition of Measurement Uncertainty
  Expression of Measurement Uncertainty
17025:2017 Requirements
  Use in selection of methods
  Use in reporting
Statistics, Units and Calculations
  Standard deviation
  Transformation of non-normal data
  Harmonisation of scales & units
Approaches to Estimation of Measurement Uncertainty
  Metrological approach
  Single Within-Laboratory Approach
Guidance Documents